Test Instrument Module System

USB Stepper Motor Controller

Model: TIMS-0201

Function Protocol

Software Specification

Review Copy

Not Released

Revision History

| ECO | Revision | Change Description | Date | Initial |
| --- | --- | --- | --- | --- |
|  | 01 | Initial Draft | 4/29/2004 | GSH |
|  | 02 | Updated codes for new SAMCOM | 11/22/04 | JED |

Table of Contents

1	Introduction	5

2	SCOPE	5

3	Contact Information	5

3.1	Sales and Support	5

4	Reference Documents	6

5	Specification	7

5.1	General	7

5.1.1	Protocol Layer-1: Physical Layer	7

5.2	Terminology	7

5.3	Control Functions	8

5.3.1	Function: USER DEFAULTS	10

5.3.2	Function: FACTORY DEFAULTS	11

5.3.3	Function: GET STATUS	12

5.3.3.1	Status Flags Table	13

5.3.4	Function: STOP	14

5.3.5	Function: SINGLE STEP	15

5.3.6	Function: CONTINUOUS STEPPING	16

5.3.7	Function: SEEK HOME	17

5.3.8	Function: SEEK LIMIT	18

5.3.9	Function: MOVE ABSOLUTE	19

5.3.10	Function: MOVE RELATIVE	20

5.3.11	Function: SET POSITION – STEP COUNT	21

5.3.12	Function: GET POSITION – STEP COUNT	22

5.3.13	Function: SET POSITION – ENCODER COUNT	23

5.3.14	Function: GET POSITION – ENCODER COUNT	24

5.3.15	Function: GET TARGET POSITION	25

5.3.16	Function: SAVE SOFT POSITION LIMITS	26

5.3.17	Function: GET SOFT POSITION LIMITS	27

5.3.18	Function: SET STEP MODE	28

5.3.19	Function: SET MOVE COUNTER MODE	29

5.3.20	Function: SET SOFT LIMITS MODE	30

5.3.21	Function: SET MOTOR CURRENT LIMIT	31

5.3.22	Function: GET MOTOR CURRENT LIMIT	32

5.3.23	Function: MEASURE MOTOR POWER	33

5.3.24	Function: SAVE NORMAL STEP RATE	34

5.3.25	Function: SET NORMAL STEP RATE	35

5.3.26	Function: GET NORMAL STEP RATE	36

5.3.27	Function: SAVE SEEK STEP RATE	37

5.3.28	Function: SET SEEK STEP RATE	38

5.3.29	Function: GET SEEK STEP RATE	39

5.3.30	Function: SAVE GPIO CONFIG	40

5.3.30.1	GPIO Configuration Table	41

5.3.31	Function: WRITE GPIO-DIO	43

5.3.32	Function: READ GPIO-DIO	44

5.3.33	Function: SET GPIO-COUNTER PERIOD	45

5.3.34	Function: READ GPIO-COUNTER	46

5.3.35	Function: SAVE GPIO-PWM FREQUENCY	47

5.3.35.1	PWM Frequency Settings	48

5.3.36	Function: SET GPIO-PWM FREQUENCY	49

5.3.37	Function: GET GPIO-PWM FREQUENCY	50

5.3.38	Function: SAVE GPIO-PWM REGISTERS	51

5.3.38.1	PWM Register Settings	52

5.3.39	Function: SET GPIO-PWM REGISTERS	53

5.3.40	Function: GET GPIO-PWM REGISTERS	54

5.3.41	Function: READ GPIO-AI	55

# Introduction

This specification provides the necessary information to design, develop, manufacture, use and maintain the WireWorks West model TIMS-0201, Stepper Motor Controller, module control software.

# SCOPE

The information presented in this document is limited to the specific software functions of the TIMS-0201.

# Contact Information

## Sales and Support

WireWorks West, Inc.

965 Mission Street, Suite 600

San Francisco, CA 94103

415-348-1400 Office

415-348-1414 Fax

415-348-1408 Technical Support

800-755-1400 Toll Free Sales

www.wireworkswest.com

# Reference Documents

| Description | Doc. No | Company/Author | Rev/Date |
| --- | --- | --- | --- |
| SW Spec, SAMCOM | 181-0001-xx | WWW/GSH |  |

# Specification

## General

This subsection will present certain consideration, rules, and design constraints specific for the development of software for use with the TIMS-0201 USB Stepper Motor Controller module.

- The TIMS-0201 utilizes a single microcontroller device with a protocol address of 0xD0.

### Protocol Layer-1: Physical Layer

The TIMS-0201 utilizes a USB serial interface device that is compliant with USB 1.1 and 2.0 full speed requirements as specified by the USB.ORG committee.

- Serial data stream characteristics: 230K Baud, 8-bit data, no parity, 1 stop bit

- The TIMS-0201 may draw up to 500 ma of 5 VDC power from the USB interface

## Terminology

|  | Term | Meaning |
| --- | --- | --- |
|  | FWD | Forward: Direction in reference to the step counter being incremented |
|  | REV | Reverse: Direction in reference to the step counter being decremented |

The step counter will be incremented by stepping in the forward direction, and decremented by stepping in the reverse direction.

The step counter will be incremented, or decremented, by one in half step mode and by two in the full step mode.

## Control Functions

TIMS-0201 Module Function Summary

|  | Function | Control Code | Comment |
| --- | --- | --- | --- |
|  | No Operation | 0x00 | Unused Code Status = Unknown Code |
|  | Load User Defaults | 0x0201 |  |
|  | Load Factory Defaults | 0x00FD |  |
|  | Get Status | 0x0200 |  |
|  | Stop | 0x0210 |  |
|  | Brake | 0x0211 | Reserved for future operation |
|  | Single Step | 0x212 |  |
|  | Continuous Stepping | 0x213 |  |
|  | Seek Home Fwd/Rev | 0x0214 |  |
|  | Seek Limit Fwd/Rev | 0x0215 |  |
|  | Move to Absolute Position | 0x0216 |  |
|  | Move to Relative Position | 0x0217 |  |
|  | Get Position (Step Count) | 0x0220 |  |
|  | Set Position (Step Count) | 0x0221 |  |
|  | Get Position (Encoder Count) | 0x0222 |  |
|  | Set Position (Encoder Count) | 0x0223 |  |
|  | Get Target Position | 0x0224 |  |
|  | Get Soft Limits | 0x0226 |  |
|  | Save Soft Limits | 0x0227 |  |
|  | Set Step Mode | 0x0228 | Half/Full Step |
|  | Set Counter Mode | 0x0229 | Encoder/Step Count |
|  | Set Soft Limits Mode | 0x022A | Enable/Disable |
|  | Get Motor Current Limit | 0x0230 |  |
|  | Set Motor Current Limit | 0x0231 |  |
|  | Measure Motor Voltage & Winding Currents | 0x0232 |  |

TIMS-0201 Module Specific Functions con’t

|  | Function | Control Code | Comment |
| --- | --- | --- | --- |
|  | Set Normal Step Rate | 0x0240 |  |
|  | Get Normal Step Rate | 0x0241 |  |
|  | Save Normal Step Rate | 0x0242 |  |
|  | Set Seek Step Rate | 0x0244 |  |
|  | Get Seek Step Rate | 0x0245 |  |
|  | Save Seek Step Rate | 0x0246 |  |
|  | Get GPIO Configuration | 0x0260 |  |
|  | Save GPIO Configuration | 0x0261 |  |
|  | Read GPIO-DIO | 0x0262 |  |
|  | Write GPIO-DIO | 0x0263 |  |
|  | Read GPIO-Counter Period | 0x0264 |  |
|  | Write GPIO-Counter | 0x0265 |  |
|  | Get PWM Frequency | 0x0268 |  |
|  | Set PWM Frequency | 0x0269 |  |
|  | Save PWM Frequency | 0x026A |  |
|  | Get PWM Registers | 0x026B |  |
|  | Set PWM Registers | 0x026C |  |
|  | Save PWM Registers | 0x026D |  |
|  | Read GPIO-ADC | 0x026F |  |

### Function: USER DEFAULTS

Function Summary

|  | Description | Restores the users defined operating settings from non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x0D |
|  | Supplied Variables | None |
|  | Returned Variables | None |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x0D |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x0D |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: FACTORY DEFAULTS

Function Summary

|  | Description | Restores factory default setting into non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x0F |
|  | Supplied Variables | None |
|  | Returned Variables | None |
|  | Comments | Factory Default Settings: Soft Limits: 2,147,483,647 Max, -2,147,483,648 Min Normal Step Rate: 200 Steps/Sec Seek Step Rate: 50 Steps/Sec GPIO Configuration: GPIO-0 thru -4 Digital In, GPIO-5 ADC 5Vref PWM Frequency: Low, 2441 Hz, 50% Duty |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x0F |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x0F |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: GET STATUS

Function Summary

|  | Description | Get controller status flags |
| --- | --- | --- |
|  | Control Code | 0x20 |
|  | Supplied Variables | None |
|  | Returned Variables | Status Flags (U16) |
|  | Comments | Refer to Status Flags table |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x20 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x20 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Status Flags, Bits[15:8] U16 – Status Flags, Bits[7:0] |

#### Status Flags Table

Status Flags Table

|  | Bit | Name | Description |
| --- | --- | --- | --- |
|  | 0 | BUSY | Stepper Motor Busy 1 = Busy, stepping operation in progress 0 = Not busy |
|  | 1 | MOVE | Move to Target Position 1 = Absolute or relative move to target position in progress 0 = Move function not in progress |
|  | 2 | SNGL | Single Step 1 = Single step in progress 0 = Move function not in progress |
|  | 3 | DIR | Step Direction 1 = Step forward, increment step counter 0 = Step reverse, decrement step counter |
|  | 4 | STEP | Step Mode 1 = Half step mode 0 = Full step mode |
|  | 5 | HOME | Seek Home 1 = Seek home position operation in progress. 0 = Seek home operation not in progress. |
|  | 6 | LIMIT | Seek Limit 1 = Seek limit position operation in progress. 0 = Seek limit operation not in progress. |
|  | 7 | COUNT | Counter Mode 1 = Use encoder counter in absolute or relative move operations. 0 = Use step counter in absolute or relative move operations. |
|  | 8 | SOFT | Soft Limits Mode 1 = Use of soft limits to stop stepping operation is enabled. 0 = Use of soft limits is disabled. |
|  | 9 | @HOME | At Home Position 1 = Home limit input is enabled and low. 0 = Home limit input is either not enabled or is high. |
|  | 10 | @FWD | At FWD Limit Position 1 = FWD limit input is enabled and low. 0 = FWD limit input is either not enabled or is high. |
|  | 11 | @REV | At REV Limit Position 1 = REV limit input is enabled and low. 0 = REV limit input is either not enabled or is high. |
|  | 12 | >MAX | Max Soft Limit Exceeded 1 = Current position exceeds the soft maximum position limit. 0 = Limit not exceeded. |
|  | 13 | @MAX | At Max Soft Limit Position 1 = Current position equals the soft maximum position limit. 0 = Limit and position not equal. |
|  | 14 | @MIN | At Min Soft Limit Position 1 = Current position equals the soft minimum position limit. 0 = Limit and position not equal. |
|  | 15 | <MIN | Min Soft Limit Exceeded 1 = Current position exceeds the soft minimum position limit. 0 = Limit not exceeded. |

### Function: STOP

Function Summary

|  | Description | Stops motor stepping operation |
| --- | --- | --- |
|  | Control Code | 0x11 |
|  | Supplied Variables | None |
|  | Returned Variables | None |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x11 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x11 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: SINGLE STEP

Function Summary

|  | Description | Performs a single step in either the forward or reverse direction |
| --- | --- | --- |
|  | Control Code | 0x13 |
|  | Supplied Variables | Direction (u8) |
|  | Returned Variables | None |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Returns an error if Busy. Returns an error if Step FWD and at FWD Limit. Returns an error if Step REV and at REV Limit. The step counter will be incremented by stepping in the forward direction, and decremented by stepping in the reverse direction. The step counter will be incremented, or decremented, by one in half step mode and by two in the full step mode. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x13 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 - Direction: 0 = REV, else FWD |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x13 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: CONTINUOUS STEPPING

Function Summary

|  | Description | Performs continuous stepping in either the forward or reverse direction |
| --- | --- | --- |
|  | Control Code | 0x14 |
|  | Supplied Variables | Direction (u8) |
|  | Returned Variables | None |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Uses Normal step rate. Returns an error if Busy. Stops if Stepping FWD and reaches FWD Limit. Stops if Stepping REV and reaches REV Limit. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x14 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 - Direction: 0 = REV, else FWD |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x14 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: SEEK HOME

Function Summary

|  | Description | Performs continuous stepping in either the forward or reverse direction until home detected |
| --- | --- | --- |
|  | Control Code | 0x15 |
|  | Supplied Variables | Direction (u8) |
|  | Returned Variables | none |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Uses Seek step rate. Returns an error if Busy. Returns an error if a DIO not configured as Home input. Reverses direction if the Fwd or Rev limit is reached. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x15 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 - Direction: 0 = REV, else FWD |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x15 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: SEEK LIMIT

Function Summary

|  | Description | Performs continuous stepping in either the forward or reverse direction until either the FWD or REV limit, as indicated by direction, is detected |
| --- | --- | --- |
|  | Control Code | 0x16 |
|  | Supplied Variables | Direction (u8) |
|  | Returned Variables | None |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Uses Seek step rate. Returns an error if Busy. Returns an error if a DIO is not configured as limit input or soft limits are not enabled. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x16 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 - Direction: 0 = REV, else FWD |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x16 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: MOVE ABSOLUTE

Function Summary

|  | Description | Initiates stepping to a specified position |
| --- | --- | --- |
|  | Control Code | 0x17 |
|  | Supplied Variables | Absolute Position (i32) |
|  | Returned Variables | Target Position (i32) |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Uses either step count or encoder count depending on Step/Encoder Count Mode. Uses Normal step rate. Returns an error if Busy. Returns an error if soft limits are enabled and absolute position is outside of the soft limits. Stops stepping if FWD or REV limit is enabled and reached. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x17 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 - Absolute Position, Bits[31:24] I32 - Absolute Position, Bits[23:16] I32 - Absolute Position, Bits[15:8] I32 - Absolute Position, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x17 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 - Target Position, Bits[31:24] I32 - Target Position, Bits[23:16] I32 - Target Position, Bits[15:8] I32 - Target Position, Bits[7:0] |

### Function: MOVE RELATIVE

Function Summary

|  | Description | Initiates stepping to a position relative to the current position |
| --- | --- | --- |
|  | Control Code | 0x18 |
|  | Supplied Variables | Relative Position (i32) |
|  | Returned Variables | Target Position (i32) |
|  | Comments | Performs either a full or half step depending on the Full/Half Step Mode. Uses either step count or encoder count depending on Step/Encoder Count Mode. Uses Normal step rate. Returns an error if Busy. Returns an error if soft limits are enabled and the computed target position is outside of the soft limits. Stops stepping if FWD or REV limit is enabled and reached. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x18 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 - Relative Position, Bits[31:24] I32 - Relative Position, Bits[23:16] I32 - Relative Position, Bits[15:8] I32 - Relative Position, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x18 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 - Target Position, Bits[31:24] I32 - Target Position, Bits[23:16] I32 - Target Position, Bits[15:8] I32 - Target Position, Bits[7:0] |

### Function: SET POSITION – STEP COUNT

Function Summary

|  | Description | Sets the step count value for the current motor position |
| --- | --- | --- |
|  | Control Code | 0x20 |
|  | Supplied Variables | New Step Count (i32) |
|  | Returned Variables | None |
|  | Comments | Step Position in half step units Returns an error if Busy. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x20 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 – New Step Count, Bits[31:24] I32 – New Step Count, Bits[23:16] I32 – New Step Count, Bits[15:8] I32 – New Step Count, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x20 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: GET POSITION – STEP COUNT

Function Summary

|  | Description | Gets the step count value for the current motor position |
| --- | --- | --- |
|  | Control Code | 0x21 |
|  | Supplied Variables | None |
|  | Returned Variables | Current Step Count (i32) |
|  | Comments | Step Position in half step units |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x21 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x21 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 – Current Step Count, Bits[31:24] I32 – Current Step Count, Bits[23:16] I32 – Current Step Count, Bits[15:8] I32 – Current Step Count, Bits[7:0] |

### Function: SET POSITION – ENCODER COUNT

Function Summary

|  | Description | Sets the encoder count value for the current motor position |
| --- | --- | --- |
|  | Control Code | 0x22 |
|  | Supplied Variables | New Encoder Count (i32) |
|  | Returned Variables | None |
|  | Comments | Returns an error if Busy. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x22 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 – New Encoder Count, Bits[31:24] I32 – New Encoder Count, Bits[23:16] I32 – New Encoder Count, Bits[15:8] I32 – New Encoder Count, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x22 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: GET POSITION – ENCODER COUNT

Function Summary

|  | Description | Gets the encoder count value for the current motor position |
| --- | --- | --- |
|  | Control Code | 0x23 |
|  | Supplied Variables | None |
|  | Returned Variables | Current Encoder Count (i32) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x23 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x23 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 – Current Encoder Count, Bits[31:24] I32 – Current Encoder Count, Bits[23:16] I32 – Current Encoder Count, Bits[15:8] I32 – Current Encoder Count, Bits[7:0] |

### Function: GET TARGET POSITION

Function Summary

|  | Description | Gets the target position value computed by the last move command |
| --- | --- | --- |
|  | Control Code | 0x24 |
|  | Supplied Variables | None |
|  | Returned Variables | Target Position (i32) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x24 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x24 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | I32 – Target Position, Bits[31:24] I32 – Target Position, Bits[23:16] I32 – Target Position, Bits[15:8] I32 – Target Position, Bits[7:0] |

### Function: SAVE SOFT POSITION LIMITS

Function Summary

|  | Description | Sets the soft FWD and REV limits for motor position, and saves the values into non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x26 |
|  | Supplied Variables | New Max Soft Position (i32) New Min Soft Position (i32) |
|  | Returned Variables | none |
|  | Comments | Returns an error if Busy. Returns an error if Length not equal to 8 Returns an error if max not greater than min. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x26 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 8 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] Data[6] Data[7] |  |  | I32 – New Max Soft Limit, Bits[31:24] I32 – New Max Soft Limit, Bits[23:16] I32 – New Max Soft Limit, Bits[15:8] I32 – New Max Soft Limit, Bits[7:0] I32 – New Min Soft Limit, Bits[31:24] I32 – New Min Soft Limit, Bits[23:16] I32 – New Min Soft Limit, Bits[15:8] I32 – New Min Soft Limit, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x23 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: GET SOFT POSITION LIMITS

Function Summary

|  | Description | Returns the soft Fwd and Rev limits for motor position |
| --- | --- | --- |
|  | Control Code | 0x27 |
|  | Supplied Variables | None |
|  | Returned Variables | Max Soft Position (i32) Min Soft Position (i32) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x27 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x27 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 8 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] Data[6] Data[7] |  |  | I32 – Max Soft Limit, Bits[31:24] I32 – Max Soft Limit, Bits[23:16] I32 – Max Soft Limit, Bits[15:8] I32 – Max Soft Limit, Bits[7:0] I32 – Min Soft Limit, Bits[31:24] I32 – Min Soft Limit, Bits[23:16] I32 – Min Soft Limit, Bits[15:8] I32 – Min Soft Limit, Bits[7:0] |

### Function: SET STEP MODE

Function Summary

|  | Description | Sets the state of the Step Mode (Half/Full) |
| --- | --- | --- |
|  | Control Code | 0x28 |
|  | Supplied Variables | New Step Mode State (u8) |
|  | Returned Variables | None |
|  | Comments | Returns an error if Busy. Note: If changing from half to full step mode and the step counter is odd, a single half step will be performed prior to any full steps as part of the next move function. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x28 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 - Step Mode: 0 = Full, else = Half |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x28 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: SET MOVE COUNTER MODE

Function Summary

|  | Description | Sets, the counter to use for move functions (Encoder/Step) |
| --- | --- | --- |
|  | Control Code | 0x29 |
|  | Supplied Variables | New Move Counter Mode State (u8) |
|  | Returned Variables | None |
|  | Comments | Returns an error if Busy. Returns an error if selecting encoder counter and a DIO has not been configured for encoder input. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x29 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – Move Counter Mode: 0 = Step, else Encoder |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x29 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: SET SOFT LIMITS MODE

Function Summary

|  | Description | Sets the soft limits mode, enable or disable the use if soft limits. |
| --- | --- | --- |
|  | Control Code | 0x2A |
|  | Supplied Variables | Soft Limits Mode (u8) |
|  | Returned Variables |  |
|  | Comments | Returns an error if Busy. If enabled, motor stepping is ceased when a limit is reached. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x2A |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – Soft Limits Mode: 0 = Disable, else Enable |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x2A |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: SET MOTOR CURRENT LIMIT

Function Summary

|  | Description | Set the motor current limit. |
| --- | --- | --- |
|  | Control Code | 0x30 |
|  | Supplied Variables | Current Limit (u8) |
|  | Returned Variables | None |
|  | Comments | Motor current limit is expressed percentage from 0% to 100%. Any value over 100% is coerced to be 100%. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x30 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – Current Limit |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x30 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: GET MOTOR CURRENT LIMIT

Function Summary

|  | Description | Returns the motor current limit. |
| --- | --- | --- |
|  | Control Code | 0x31 |
|  | Supplied Variables | None |
|  | Returned Variables | Current Limit (u8) |
|  | Comments | Motor current limit is expressed percentage from 0% to 100%. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x31 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x31 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – Current Limit |

### Function: MEASURE MOTOR POWER

Function Summary

|  | Description | Returns the motor source voltage and winding currents. |
| --- | --- | --- |
|  | Control Code | 0x32 |
|  | Supplied Variables | None |
|  | Returned Variables | Motor Source Voltage in millivolts (u16) Motor Winding-A Current in milliamps (u16) Motor Winding-B Current in milliamps (u16) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x32 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x32 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 6 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] |  |  | U16 – Motor Voltage, Bits[15:8] U16 – Motor Voltage, Bits[7:0] U16 – Winding-A Current, Bits[15:8] U16 – Winding-A Current, Bits[7:0] U16 – Winding-B Current, Bits[15:8] U16 – Winding-B Current, Bits[7:0] |

### Function: SAVE NORMAL STEP RATE

Function Summary

|  | Description | Set the step rate for normal stepping operations, and saves the setting into non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x40 |
|  | Supplied Variables | Steps/Sec (u16) - Range: 4000 Max, 2 Min |
|  | Returned Variables | None |
|  | Comments | An error is returned if the setting is out of range. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x40 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x40 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: SET NORMAL STEP RATE

Function Summary

|  | Description | Set the step rate for normal stepping operations |
| --- | --- | --- |
|  | Control Code | 0x41 |
|  | Supplied Variables | Steps/Sec (u16) - Range: 4000 Max, 2 Min |
|  | Returned Variables | None |
|  | Comments | An error is returned if the setting is out of range. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x41 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x41 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: GET NORMAL STEP RATE

Function Summary

|  | Description | Return the step rate for normal stepping operations |
| --- | --- | --- |
|  | Control Code | 0x43 |
|  | Supplied Variables | None |
|  | Returned Variables | Steps/Sec (u16) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x43 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x43 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 3 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

### Function: SAVE SEEK STEP RATE

Function Summary

|  | Description | Set the step rate for seek stepping operations, and saves the setting into non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x44 |
|  | Supplied Variables | Steps/Sec (u16) - Range: 4000 Max, 2 Min |
|  | Returned Variables | None |
|  | Comments | An error is returned if the setting is out of range. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x44 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x44 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: SET SEEK STEP RATE

Function Summary

|  | Description | Set the step rate for seek stepping operations. |
| --- | --- | --- |
|  | Control Code | 0x45 |
|  | Supplied Variables | Steps/Sec (u16) - Range: 4000 Max, 2 Min |
|  | Returned Variables | None |
|  | Comments | An error is returned if the setting is out of range. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x45 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x45 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

### Function: GET SEEK STEP RATE

Function Summary

|  | Description | Return the step rate for seek stepping operations |
| --- | --- | --- |
|  | Control Code | 0x46 |
|  | Supplied Variables | None |
|  | Returned Variables | Steps/Sec (u16) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x46 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x46 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 3 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – Steps/Sec, Bits[15:8] U16 – Steps/Sec, Bits[7:0] |

### Function: SAVE GPIO CONFIG

Function Summary

|  | Description | Sets the General Purpose IO pin configurations, and saves settings into non-volatile memory |
| --- | --- | --- |
|  | Control Code | 0x60 |
|  | Supplied Variables | GPIO-0 Config (u8) GPIO-1 Config (u8) GPIO-2 Config (u8) GPIO-3 Config (u8) GPIO-4 Config (u8) GPIO-5 Config (u8) |
|  | Returned Variables | None |
|  | Comments | Returns an error if Busy. Refer to the GPIO Configuration Table |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x60 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 6 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] |  |  | U8 – GPIO-0 Config U8 – GPIO-1 Config U8 – GPIO-2 Config U8 – GPIO-3 Config U8 – GPIO-4 Config U8 – GPIO-5 Config |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x60 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

#### GPIO Configuration Table

GPIO Configuration Table

|  | GPIO | Configuration |
| --- | --- | --- |
|  | 0 | GPIO-0 0 = Digital Input 1 = Digital Output 2 = Home Limit Input 3 = FWD Limit Input 4 = REV Limit Input 5 = Quadrature Encoder Phase-A Input 6 = Quadrature Encoder Phase-B Input |
|  | 1 | GPIO-1 0 = Digital Input 1 = Digital Output 2 = Home Limit Input 3 = FWD Limit Input 4 = REV Limit Input 5 = Quadrature Encoder Phase-A Input 6 = Quadrature Encoder Phase-B Input |
|  | 2 | GPIO-2 0 = Digital Input 1 = Digital Output 2 = Home Limit Input 3 = FWD Limit Input 4 = REV Limit Input |
|  | 3 | GPIO-3 0 = Digital Input 1 = Digital Output 2 = Home Limit Input 3 = FWD Limit Input 4 = REV Limit Input 5 = Counter/Timer Input |
|  | 4 | GPIO-4 0 = Digital Input 1 = Digital Output 2 = Home Limit Input 3 = FWD Limit Input 4 = REV Limit Input 5 = PWM Output |
|  | 5 | GPIO-5 0 = Analog Input to 10-Bit ADC, 4.096 VDC Reference. 1 = Analog Input to 10-Bit ADC, USB 5 VDC Reference. |

### Function: GET GPIO CONFIG

Function Summary

|  | Description | Sets the General Purpose IO pin configurations, and saves settings into non-volatile memory |
| --- | --- | --- |
|  | Control Code | 0x61 |
|  | Supplied Variables | None |
|  | Returned Variables | GPIO-0 Config (u8) GPIO-1 Config (u8) GPIO-2 Config (u8) GPIO-3 Config (u8) GPIO-4 Config (u8) GPIO-5 Config (u8) |
|  | Comments | Refer to the GPIO Configuration Table |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x61 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x61 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 6 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] |  |  | U8 – GPIO-0 Config U8 – GPIO-1 Config U8 – GPIO-2 Config U8 – GPIO-3 Config U8 – GPIO-4 Config U8 – GPIO-5 Config |

### Function: WRITE GPIO-DIO

Function Summary

|  | Description | Write GPIO-DO output pin logic states |
| --- | --- | --- |
|  | Control Code | 0x62 |
|  | Supplied Variables | GPIO-DIO (u8) |
|  | Returned Variables | None |
|  | Comments | U8 bits 0 through 4 maps to GPIO-DO pin 0 through 4 respectively. Only those pins configured as outputs will be effected. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x62 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – GPIO-DO |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x62 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: READ GPIO-DIO

Function Summary

|  | Description | Read GPIO-DIO pin logic states |
| --- | --- | --- |
|  | Control Code | 0x63 |
|  | Supplied Variables | None |
|  | Returned Variables | GPIO-DIO (u8) |
|  | Comments | U8 bits 0 through 4 maps to GPIO-DIO pins 0 through 4 respectively. Both input and output states are read. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x63 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x63 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – GPIO-DIO |

### Function: SET GPIO-COUNTER PERIOD

Function Summary

|  | Description | Set the counter period |
| --- | --- | --- |
|  | Control Code | 0x64 |
|  | Supplied Variables | Count Period in milliseconds (u8) |
|  | Returned Variables | None |
|  | Comments | Returns an error if input not configured for counter operation. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) | Non-zero, (tbd) by sender |
|  | REF | U8 | (tbd) | (tbd) by sender |
|  | CONTROL | U8 | 0x64 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 1 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] |  |  | U8 – Count Period; 0 = OFF, 255 = Always Counting, else N Milliseconds |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x64 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: READ GPIO-COUNTER

Function Summary

|  | Description | Returns GPIO Counter information |
| --- | --- | --- |
|  | Control Code | 0x65 |
|  | Supplied Variables | None |
|  | Returned Variables | Period to go (u8) Count Period (u8) Counter (u32) Overflow (u8) |
|  | Comments | Returns an error if input not configured for counter operation. If the Period To Go is 0 and the Count Period is >0 and <255 then the period to go will be refreshed with the count period and the counter will be cleared at the end of this function. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x65 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x65 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 7 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] Data[4] Data[5] Data[6] |  |  | U8 – Period To Go U8 – Count Period U32 – Counter, Bits[31:24] U32 – Counter, Bits[23:16] U32 – Counter, Bits[15:8] U32 – Counter, Bits[7:0] U8 – Counter Overflow; 1 = Overflow, 0 = OK |

### Function: SAVE GPIO-PWM FREQUENCY

Function Summary

|  | Description | Set the PWM frequency and duty cycle, and save setting into non-volatile memory. Actual PWM register settings computed are also saved into non-volatile memory. |
| --- | --- | --- |
|  | Control Code | 0x68 |
|  | Supplied Variables | Range (U8) Frequency (U16) Duty Cycle (U8) |
|  | Returned Variables | None |
|  | Comments | See PWM Frequency Settings This function alters the current PWM Register function settings. This function is intended to provide a rapid means to set PWM output frequency and duty cycle. For higher accuracy control see the PWM Register functions. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x68 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Range U16 – Frequency, Bits[15:8] U16 – Frequency, Bits[7:0] U8 – Duty Cycle |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x68 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

#### PWM Frequency Settings

PWM Frequency Range

0 = High Range: 10 MHz – 39 KHz, frequency expressed in KHz

1 = Med Range: 2.5 MHz – 9 KHz, frequency expressed in KHz

2 = Low Range: 625 KHz – 2441 Hz, frequency expressed in Hz

Duty Cycle

0 – 100%

As the frequency in a given range increases, the accuracy of the duty cycle decreases.  At the highest frequency of each range the resolution of the duty cycle is 4 steps; 100% (always outputs a 1), 75%, 50%, 25%, and 0% (always outputs a 0).

### Function: SET GPIO-PWM FREQUENCY

Function Summary

|  | Description | Set the PWM frequency and duty cycle. |
| --- | --- | --- |
|  | Control Code | 0x69 |
|  | Supplied Variables | Range (U8) Frequency (U16) Duty Cycle (U8) |
|  | Returned Variables | None |
|  | Comments | See PWM Frequency Settings This function alters the current PWM Register function settings. This function is intended to provide a rapid means to set PWM output frequency and duty cycle. For higher accuracy control see the PWM Register functions. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x69 |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Range U16 – Frequency, Bits[15:8] U16 – Frequency, Bits[7:0] U8 – Duty Cycle |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x69 |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: GET GPIO-PWM FREQUENCY

Function Summary

|  | Description | Returns the PWM frequency and duty cycle. |
| --- | --- | --- |
|  | Control Code | 0x6A |
|  | Supplied Variables | None |
|  | Returned Variables | Range (U8) Frequency (U16) Duty Cycle (U8) |
|  | Comments | See PWM Frequency Table The PWM Register functions do not alter these settings, therefore if PWM Register functions have been used, these settings will not reflect the actual output condition. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x6A |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x6A |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Range U16 – Frequency, Bits[15:8] U16 – Frequency, Bits[7:0] U8 – Duty Cycle |

### Function: SAVE GPIO-PWM REGISTERS

Function Summary

|  | Description | Set the PWM registers, and save settings into non-volatile memory |
| --- | --- | --- |
|  | Control Code | 0x6B |
|  | Supplied Variables | Prescale (U8) PWM Period (U8) Duty Period (U16) |
|  | Returned Variables | None |
|  | Comments | See PWM Register Settings This function does not change any PWM Frequency function settings. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x6B |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Prescale U8 – PWM Period U16 – Duty Period, Bits[15:8] U16 – Duty Period, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x6B |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

#### PWM Register Settings

PWM Prescale

0 = Base frequency/period = 10 MHz/100ns

1 = Base frequency/period = 2.5 MHz/250ns

2 = Base frequency/period = 625 KHz/1.6us

PWM Period

Period = (PWM Period+1) * 4 * (Base Period) * Prescale

Frequency = 1/Period

PWM Duty Period

10-Bit max resolution

Period = (PWM Duty Period+1) * (Base Period) * Prescale

Note:  A Duty period greater then the PWM period will result in the output always being a 1.

The effective duty period value is in the range of zero to (PWM Period +1 *4), but less than 0x0400.

PWM Resolution (max) = [log(10 MHz/Fpwm)]/log(2) bits

### Function: SET GPIO-PWM REGISTERS

Function Summary

|  | Description | Set the PWM registers. |
| --- | --- | --- |
|  | Control Code | 0x6C |
|  | Supplied Variables | Prescale (U8) PWM Period (U8) Duty Period (U16) |
|  | Returned Variables | None |
|  | Comments | See PWM Register Settings This function does not change any PWM Frequency function settings. |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x6C |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Prescale U8 – PWM Period U16 – Duty Period, Bits[15:8] U16 – Duty Period, Bits[7:0] |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x6C |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

### Function: GET GPIO-PWM REGISTERS

Function Summary

|  | Description | Return the PWM register settings. |
| --- | --- | --- |
|  | Control Code | 0x6D |
|  | Supplied Variables | None |
|  | Returned Variables | Prescale (U8) PWM Period (U8) Duty Period (U16) |
|  | Comments | See PWM Register Settings |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x6D |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | none |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x6D |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 4 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] Data[2] Data[3] |  |  | U8 – Prescale U8 – PWM Period U16 – Duty Period, Bits[15:8] U16 – Duty Period, Bits[7:0] |

### Function: READ GPIO-AI

Function Summary

|  | Description | Reads the GPIO analog input |
| --- | --- | --- |
|  | Control Code | 0x6F |
|  | Supplied Variables | None |
|  | Returned Variables | GPIO-AI in millivolts (u16) |
|  | Comments |  |

Command Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | 0xD0 |  |
|  | FROM | U8 | (tbd) |  |
|  | REF | U8 | (tbd) |  |
|  | CONTROL | U8 | 0x6F |  |
|  | STATUS | U8 | (xxx) | Don’t care, ignored by receiver |
|  | Data_Length | U8 | 0 |  |
|  | DATA | [u8] | None |  |

Response Packet

|  | Element | Type | Value | Description |
| --- | --- | --- | --- | --- |
|  | TO | U8 | (tbd) | Echo of Command Packet FROM |
|  | FROM | U8 | 0xD0 |  |
|  | REF | U8 | (tbd) | Echo of Command Packet REF |
|  | CONTROL | U8 | 0x6F |  |
|  | STATUS | U8 | (tbd) |  |
|  | Data_Length | U8 | 2 |  |
|  | DATA | [u8] |  | See content below |
|  | Data[0] Data[1] |  |  | U16 – GPIO-AI, Bits[15:8] U16 – GPIO-AI, Bits[7:0] |